Earth and Space Science

Supporting Information for

Improving ECC Ozonesonde Data Quality: Assessment of Current Methods and Outstanding Issues

David W. Tarasick1, Herman G.J. Smit2, Anne M. Thompson3 , Gary A. Morris4, J.C. Witte5, J. Davies1, Tatsumi Nakano6, Roeland van Malderen7, Ryan M. Stauffer8, Bryan J. Johnson9, Rene Stübi10, Samuel J. Oltmans9 and Holger Vömel11

#### 1Environment and Climate Change Canada, 4905 Dufferin Street, Downsview, ON, M3H 5T4 Canada

#### 2Institute for Energy and Climate Research: Troposphere (IEK-8), Research Centre Juelich (FZJ), Juelich, Germany.

3NASA Goddard Space Flight Center, Greenbelt, Maryland, USA

4St. Edward’s University, Austin, TX, USA

5Earth Observing Laboratory, National Center for Atmospheric Research, Boulder, Colorado, USA

6Japan Meteorological Agency, Tokyo, Japan

7Royal Meteorological Institute of Belgium, Brussels, Belgium

8NASA Goddard Space Flight Center and University of Maryland Earth System Science Interdisciplinary Center, Maryland, USA

9NOAA/ESRL Global Monitoring Division, Boulder, Colorado, USA

10MeteoSwiss Aerological Station, Federal Office of Meteorology and Climatology MeteoSwiss, Payerne, Switzerland

11National Center for Atmospheric Research, Boulder, Colorado, USA

Corresponding author: David W. Tarasick (david.tarasick@canada.ca)

Contents of this file

Text S1.  WCCOS and JOSIE

Text S2.  Pump Temperature Corrections

Text S3.  Water vapour saturation pressure

Table S1.  Water vapour saturation pressure over liquid water determined after the empirical formula by Magnus-Teten (Murray, 1967)

Table S2.  Pump flow efficiencies (ηP) as a function of air pressure for ECC-ozone sondes reported by (i) Komhyr, 1986; (ii) Komhyr et al., 1995; (iii) Johnson et al. 2002: GMD & UWYO; (iv) Nakano et al., 2017: JMA.

Table S3.  Mean and standard deviations of pressure offsets as a function of altitude and model.  The “N” values are for all sondes, those that reach at least 26 km, and those that reach at least 30 km. Data from Stauffer et al. (2014).

Text S4.  Cathode Sensing Solution Types

Table S4.  Most common sensing solution types (SST) and their chemical composition in aqueous solution for cathode cells of ECC-ozonesondes. For the anode sensing solution, a KI-saturated cathode solution is employed.

Text S1.  WCCOS and JOSIE

The environmental simulation chamber at the Forschungszentrum Jülich, Germany was established as the World Calibration Centre for Ozone Sondes (WCCOS) in 1996. The heart of the facility is a temperature-controlled vacuum chamber with a test room volume of about 500 liters (80 x 80 x 80 cm). Within the chamber pressure and temperature can be dynamically regulated, with pressures between 5 and 1000 hPa and temperatures between 200 and 300 K, with a maximum rate of ±2K/min. Temperature variations of the air as well as the wall inside the test chamber can be maintained within ±0.2 K. Pressure, temperature, and ozone concentration can be varied rapidly, enabling simulation of flight conditions for ozone soundings up to an altitude of 35 km (Smit et al., 2000).

The ozone reference is a fast response dual-beam UV absorption photometer (OPM), originally developed by Proffitt and McLaughlin (1983) for use on stratospheric balloons. The instrument was flown during Balloon Ozone Intercomparison (BOIC) missions in 1983/1984 (Hilsenrath et al., 1986) and on the Balloon Experiment on Standards for Ozone Sondes (BESOS) field campaign in Wyoming, in 2004 (Deshler et al., 2008). The OPM is an absolute measuring device with a 1-s response time at a sampling volume flow rate of about 8 l/min. The overall accuracy of ozone measurements made by the OPM is better than ±2% for simulated altitudes up to 25 km (pressures down to 25 hPa) and ±3.5 % at 30-35 km altitude (12-5 hPa). The instrument resides in a separate vacuum vessel which is connected to the main environmental chamber such that the UV-photometer has the same pressure conditions as inside the test chamber.

The initial Jülich Ozone Sonde Intercomparison Experiment (JOSIE), performed in 1996 (Smit and Kley, 1998), was directed toward implementing a global quality assurance plan for ozonesondes in routine use, and followed a series of field intercomparison campaigns (Attmannspacher and Dütsch, 1970; 1981; Hilsenrath et al., 1986; Kerr et al, 1994). Subsequent experiments have examined manufacturing differences and sensing solution types (Smit and Straeter, 2004; (Smit et al., 2007), Standard Operating Procedures (SOPs) (Smit and ASOPOS panel, 2014), homogenization and uncertainties (Smit et al., 2012), and special issues in the tropics (Thompson et al., 2019).  JOSIE experiments have provided, and continue to provide, essential data to improve ozonesonde data quality and to ensure that the global ozonesonde network continues to supply reliable data to the larger atmospheric research and remote sensing communities.

Text S2.  Pump Temperature Corrections

Since the mid-1990s all ECC sondes measure pump temperature via a thermistor mounted in a hole drilled in the pump body (TP,Int). In previous decades temperature measurements were made with thermistors at other locations on the instrument. Smit et al. (2012) have developed correction formulas for these older temperature measurements, relating them to the standard measurement TP,Int, with estimates of the additional uncertainty incurred in each case. In all cases PAir is in hPa.

Case I: Box temperature measurements by rod thermistor in analog sounding systems:

PAir  ≥ 40 hPa	         (S-1)

6 < PAir < 40 hPa           (S-2)

PAir  ≤ 6 hPa	         (S-3)

The uncertainty of this parameterization is estimated to be ±1 K.

Case II: Box temperature measurements by thermistor in digital sonde systems:

The thermistor position was variable, but suspended in the ozonesonde box near the pump. Typically the temperature difference follows that of Case III.

Case III: External pump (taped thermistor) measurements in digital sonde systems:

PAir  > 70 hPa         (S-4)

15 ≤  PAir ≤ 70 hPa       (S-5)

5 ≤ PAir  ≤ 15 hPa       (S-6)

The uncertainty of this correction is estimated as:

PAir  > 70 hPa          (S-7)

PAir ≤ 70 hPa           (S-8)

Case IV: External pump (epoxied/glued thermistor) measurements in digital sonde systems:

PAir  > 40 hPa	       (S-9)

3 ≤ PAir ≤ 40 hPa         (S-10)

The uncertainty of this parameterization is estimated to be ±0.5 K.

Text S3.  Water vapour saturation pressure

The saturated water vapor pressure at laboratory temperature TLab can be calculated by way of  several empirical formulas. We use the Magnus-Teten formula (Murray, 1967):

(S-11)

where TLab is in ºC, and PSat,H2O is found from Table S1.

Table S1.  Water vapour saturation pressure over liquid water determined after the empirical formula by Magnus-Teten (Murray, 1967)

| Temperature (oC) | PSat,H2O (hPa) After Magnus-Teten (Murray, 1967) |
| --- | --- |
| 0 | 6,1 |
| 5 | 8,7 |
| 10 | 12,3 |
| 15 | 17,1 |
| 20 | 23,4 |
| 25 | 31,7 |
| 30 | 42,4 |
| 35 | 56,2 |
| 40 | 73,7 |
| 45 | 95,8 |
| 50 | 123,3 |

Table S2.  Pump flow efficiencies (ηP) as a function of air pressure for ECC-ozone sondes reported by (i) Komhyr, 1986; (ii) Komhyr et al., 1995; (iii) Johnson et al. 2002: GMD & UWYO; (iv) Nakano et al., 2017: JMA.

| Pressure [hPa] | ECC (SPC-6a) Komhyr, 1986 | ECC (EN-SCI) Komhyr et al.,1995 | ECC (GMD) Johnson et al.,2002 | ECC (UWYO) Johnson et al.,2002 | ECC (JMA) Nakano, 2016 |
| --- | --- | --- | --- | --- | --- |
| 1000<>200 | 1 | 1 | 1 | 1 | 1 |
| 100 | 1.011 ± 0.005 | 1.007 ± 0.005 | 1.033 ± 0.011 | 1.022 ± 0.011 | 1.025 ± 0.008 |
| 50 | 1.015 ± 0.006 | 1.018 ± 0.005 | 1.052 ± 0.012 | 1.037 ± 0.013 | 1.039 ± 0.010 |
| 30 | 1.022 ± 0.008 | 1.029 ± 0.008 | 1.070 ± 0.015 | 1.049 ± 0.016 | 1.055 ± 0.012 |
| 20 | 1.032 ± 0.009 | 1.041 ± 0.012 | 1.089 ± 0.018 | 1.066 ± 0.020 | 1.073 ± 0.013 |
| 10 | 1.055 ± 0.010 | 1.066 ± 0.023 | 1.145 ± 0.020 | 1.120 ± 0.032 | 1.122 ± 0.016 |
| 7 | 1.070 ± 0.012 | 1.087 ± 0.024 | 1.195 ± 0.025 | 1.166 ± 0.040 | 1.162 ± 0.019 |
| 5 | 1.092 ± 0.014 | 1.125 ± 0.024 | 1.260 ± 0.037 | 1.224 ± 0.051 | 1.213 ± 0.023 |

Table S3.  Mean and standard deviations of pressure offsets as a function of altitude and model.  The “N” values are for all sondes, those that reach at least 26 km, and those that reach at least 30 km. Data from Stauffer et al. (2014).

| Type | iMet |  | iMet-P |  | iMet-S |  | RS80 |  | RS92 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| N (all, 26, 30) | 106, 61, 37 |  | 52, 46, 25 |  | 69, 53, 42 |  | 247, 162, 91 |  | 257, 181, 144 |  |
| Alt (km) | <p> (hPa) |  (hPa) | <p> (hPa) |  (hPa) | <p> (hPa) |  (hPa) | <p> (hPa) |  (hPa) | <p> (hPa) |  (hPa) |
| 0 |  |  |  |  |  |  |  |  |  |  |
| 1 | 0.08 | 2.21 | -0.45 | 1.11 | 0.52 | 1.39 | -0.34 | 1.57 | 1.11 | 1.42 |
| 2 | -0.04 | 1.92 | -0.38 | 1.13 | 0.44 | 1.35 | 0.05 | 1.56 | 0.56 | 1.41 |
| 3 | 0.00 | 1.90 | -0.34 | 1.19 | 0.35 | 1.07 | -0.17 | 1.46 | 0.35 | 1.17 |
| 4 | -0.04 | 1.88 | -0.32 | 1.14 | 0.27 | 1.09 | -0.47 | 1.43 | 0.23 | 1.03 |
| 5 | -0.19 | 1.79 | -0.33 | 1.04 | 0.24 | 1.22 | -0.61 | 1.33 | 0.07 | 0.92 |
| 6 | -0.24 | 1.73 | -0.39 | 1.06 | 0.16 | 1.31 | -0.70 | 1.26 | -0.03 | 0.85 |
| 7 | -0.35 | 1.70 | -0.43 | 1.05 | 0.17 | 0.65 | -0.83 | 1.26 | -0.11 | 0.78 |
| 8 | -0.45 | 1.68 | -0.78 | 1.61 | 0.11 | 0.63 | -0.85 | 1.21 | -0.16 | 0.75 |
| 9 | -0.57 | 1.61 | -0.87 | 1.49 | 0.05 | 0.60 | -0.91 | 1.17 | -0.21 | 0.70 |
| 10 | -0.62 | 1.55 | -0.98 | 1.45 | -0.01 | 0.58 | -0.94 | 1.14 | -0.26 | 0.65 |
| 11 | -0.66 | 1.56 | -1.16 | 1.42 | -0.10 | 0.54 | -0.97 | 1.15 | -0.27 | 0.62 |
| 12 | -0.74 | 1.54 | -1.26 | 1.41 | -0.16 | 0.58 | -0.98 | 1.16 | -0.27 | 0.58 |
| 13 | -0.80 | 1.58 | -1.41 | 1.40 | -0.23 | 0.68 | -1.02 | 1.16 | -0.24 | 0.75 |
| 14 | -0.87 | 1.64 | -1.52 | 1.38 | -0.26 | 0.73 | -1.02 | 1.17 | -0.28 | 0.50 |
| 15 | -0.91 | 1.66 | -1.66 | 1.38 | -0.28 | 0.72 | -1.01 | 1.29 | -0.27 | 0.47 |
| 16 | -0.99 | 1.54 | -1.56 | 0.72 | -0.27 | 0.77 | -0.94 | 1.28 | -0.25 | 0.45 |
| 17 | -1.05 | 1.53 | -1.83 | 1.36 | -0.31 | 0.84 | -0.94 | 1.20 | -0.24 | 0.42 |
| 18 | -1.07 | 1.53 | -1.87 | 1.34 | -0.35 | 0.88 | -1.01 | 1.29 | -0.22 | 0.40 |
| 19 | -1.12 | 1.54 | -1.94 | 1.27 | -0.43 | 0.94 | -0.95 | 1.41 | -0.20 | 0.35 |
| 20 | -1.05 | 1.32 | -1.93 | 1.20 | -0.44 | 0.90 | -1.00 | 1.36 | -0.19 | 0.34 |
| 21 | -1.09 | 1.34 | -1.97 | 1.12 | -0.49 | 0.90 | -0.96 | 1.23 | -0.17 | 0.31 |
| 22 | -1.08 | 1.34 | -1.96 | 1.07 | -0.53 | 0.91 | -0.98 | 1.31 | -0.15 | 0.30 |
| 23 | -1.02 | 1.26 | -1.96 | 0.99 | -0.58 | 0.92 | -0.99 | 1.30 | -0.14 | 0.28 |
| 24 | -0.91 | 0.89 | -1.95 | 0.97 | -0.61 | 0.94 | -1.00 | 1.39 | -0.14 | 0.27 |
| 25 | -0.92 | 0.82 | -1.94 | 0.89 | -0.59 | 0.94 | -0.99 | 1.33 | -0.13 | 0.27 |
| 26 | -0.93 | 0.83 | -1.95 | 0.85 | -0.58 | 0.95 | -1.01 | 1.35 | -0.12 | 0.27 |
| 27 | -0.94 | 0.87 | -1.94 | 0.81 | -0.58 | 0.94 | -1.00 | 1.38 | -0.12 | 0.27 |
| 28 | -0.95 | 0.93 | -1.99 | 0.76 | -0.59 | 0.95 | -1.06 | 1.40 | -0.12 | 0.26 |
| 29 | -0.99 | 0.81 | -1.85 | 0.70 | -0.66 | 1.01 | -1.01 | 1.44 | -0.11 | 0.26 |
| 30 | -1.23 | 0.77 | -1.89 | 0.76 | -0.70 | 1.03 | -1.04 | 1.60 | -0.11 | 0.25 |
| > 26 km | -1.01 | 0.84 | -1.92 | 0.78 | -0.62 | 0.98 | -1.02 | 1.43 | -0.12 | 0.26 |
| > 20 km | -1.07 | 1.30 | -1.95 | 0.89 | -0.51 | 0.89 | -0.97 | 1.27 | -0.12 | 0.26 |

Text S4.  Cathode Sensing Solution Types

Table S1 lists the most commonly used cathode sensing solution types (SST).  SST1.0 (1.0% KI and full pH-buffer) is the conventional sensing solution based on the guidelines described by Komhyr (1986) and is most widely used for the ozonesonde types SPC-6A manufactured by Science Pump Corporation (1996).  SST0.5 (0.5% KI and half pH-buffer) is the sensing solution recommended by ENSCI Corporation (1996) since 1996 as the cathode solution for their ENSCI-type of ozonesondes.  Both combinations of SPC-6A and SST1.0 and ENSCI-Z and SST0.5 are also recommended for the global ozonesonde networks by the WMO/GAW (Smit and ASOPOS panel, 2014), respectively. From 1996 - 2007 NOAA/ESRL ozonesonde sites used a third sensing solution type, SST2.0, which is a non-pH-buffered solution with no KBr, but with 2.0% KI for the cathode cell (Johnson et al, 2002), and since 2007 a 1.0% KI and 1/10th buffer solution SST0.1 (Sterling et al., 2018).

Table S4.  Most common sensing solution types (SST) and their chemical composition in aqueous solution for cathode cells of ECC-ozonesondes. For the anode sensing solution, a KI-saturated cathode solution is employed.

(a) Komhyr (1986) , SPC-6A instruction manual (Science Pump Corporation, 1996)

(b) ENSCI-Z instruction manual (EN-SCI Corporation, 1996)

(c) Johnson et al. (2002)

(d) Johnson, private communication (2007)