Supplementary Tables

Table S1 Plasma bonding conditions

| Gas | Oxygen | Set the gauge on tank to 15 psi |
| --- | --- | --- |
| FWD/REF | 15W/0W |  |
| mTorr Vacuum | set between 250 - 270 mTorr |  |
| RECIPE PARAMETERS |  |  |
| PDMS ratio | 10:1 (base to curing agent) |  |
| RF-Vacuum | 250mTorr |  |
| REF | 0W |  |
| FWD | 15W |  |
| Time | 30 seconds |  |
| Purge time | 10 seconds |  |
| Vacuum set point | 101mTorr |  |

Table S2 Nominal feature dimension vs. measured feature dimensions through SEM

| Feature | Nominal dimension | Measured dimension |
| --- | --- | --- |
| 1st Single cell outlet | 3 µm | 2.55 µm |
| 2nd Single cell outlet (chip 1) | 3 µm | 2.60 µm |
| Pillar width | 3 µm | 2.66 µm |
| Pillar broadness | 5 µm | 4.67 µm |

Table S3 Feature height measurement through profilometry for a given number of DRIE cycles

| DRIE cycle | Duration [min] | Channel depth |
| --- | --- | --- |
| 14 | 3.54 | 10.84 µm |
| 15 | 4.1 | 11.67 µm |
| 23 | 6.18 | 16.5 µm |